AyuScan

AI-Powered Real-Time Wearable Health Monitoring System

Technical Documentation  |  v1.0  |  August 2026

AyuScan is a non-invasive, multi-modal wearable biomedical device that continuously acquires ECG, Photoplethysmography (PPG), blood oxygen saturation (SpO₂), and body temperature signals via an ESP32-C6 microcontroller. Signal data is streamed over Bluetooth Low Energy (BLE) to a Python gateway where three purpose-built AI models perform real-time cardiovascular condition classification, blood pressure estimation, and SpO₂ trend prediction. Results are broadcast via WebSocket to a live React dashboard for clinical visualisation.

# Table of Contents

1	System Overview

2	Hardware Architecture & Bone Conduction Sensing Technology

2.1	Seeed Studio XIAO ESP32-C6

2.2	BioAmp EXG Pill  –  ECG Acquisition

2.3	MAX30102  –  PPG & SpO₂ Sensing

2.4	MLX90614  –  Infrared Body Temperature

2.5	Bone Conduction & Contact-Based PPG Principle

3	Firmware  –  ayuScan_direct.ino

3.1	BLE Communication Protocol

3.2	Sensor Initialisation & I²C Bus Management

3.3	Timer-Driven ECG Acquisition at 250 Hz

3.4	PPG FIFO Streaming at 100 Hz

4	Backend  –  Python Gateway (ble_server.py)

4.1	BLE-to-WebSocket Bridge

4.2	PPG Signal Processing Algorithm

4.3	Data Logging & Persistence

5	AI Model 1  –  ECG Cardiac Condition Classifier  (1D-CNN)

5.1	Model Architecture

5.2	Dataset  –  PTB-XL

5.3	Training Procedure

5.4	Inference Pipeline

6	AI Model 2  –  Blood Pressure Estimator  (XGBoost)

6.1	Feature Engineering

6.2	Dataset  –  VitalDB

6.3	Model Details & Metrics

7	AI Model 3  –  SpO₂ Trend Classifier  (GRU / 1D-CNN)

7.1	Model Architecture

7.2	Dataset  –  BIDMC Multiparameter Dataset

7.3	Label Schema & Inference

8	Dashboard  –  React Frontend

9	System Data Flow

10	Dependencies & Environment

11	Deployment Guide

12	Clinical Safety Notes

13	Glossary

# 1  System Overview

AyuScan is designed to address the critical shortage of continuous, affordable biomedical monitoring in clinical and home settings. It combines consumer-grade sensors with production-quality AI inference to deliver a full cardiovascular screening experience from a single wrist-worn or finger-clip device.

Core capabilities at a glance:

ECG Streaming: 250 Hz single-lead ECG via BioAmp EXG Pill, streamed over BLE in 100 ms JSON batches.

Heart Rate & HRV: Derived from PPG using bandpass filtering + scipy peak detection. Reports BPM, SDNN, RMSSD.

SpO₂: Computed using the Maxim ratio-of-ratios method on IR and Red photoplethysmograph channels.

Body Temperature: Non-contact infrared measurement via MLX90614 at 1 Hz.

Blood Pressure (cuffless): Estimated from Pulse Transit Time (PTT) + HRV features using dual XGBoost regressors.

ECG Cardiac Classification: 4-class (Normal, Bradycardia, Tachycardia, Arrhythmia) using a 1D-CNN trained on PTB-XL.

SpO₂ Trend Forecasting: 4-class trend prediction (Stable → Critical) using a GRU network trained on BIDMC dataset.

Live Dashboard: React + Vite SPA connected via WebSocket displaying real-time waveforms and AI predictions.

🩺 Medical Disclaimer AyuScan is a research and educational prototype. It is NOT a certified medical device. All AI inferences are supplemental indicators and must NOT replace professional clinical judgement. Blood Pressure estimation via PTT is affected by individual physiology and calibration.

# 2  Hardware Architecture & Bone Conduction Sensing Technology

AyuScan's hardware stack integrates five key components on a shared 400 kHz I²C bus (with automatic clock-stepping to 100 kHz for SMBus-compatible sensors). All components communicate with the Seeed Studio XIAO ESP32-C6 over I²C (GPIO 4/SDA, GPIO 5/SCL) or directly via the ADC pin (GPIO 0 for ECG).

## 2.1  Seeed Studio XIAO ESP32-C6

The XIAO ESP32-C6 is the computational core of AyuScan. It runs the Arduino-compatible firmware that drives all sensors, manages the BLE GATT server, executes hardware timer interrupts for ECG sampling, and encodes all data as JSON packets.

| Parameter | Value |
| --- | --- |
| SoC | Espressif ESP32-C6 (RISC-V 160 MHz) |
| BLE | Bluetooth 5.0 LE ¦ GATT Server with 2 Notify Characteristics |
| MTU | 512 bytes (configured to prevent JSON truncation) |
| I²C | GPIO 4 (SDA) ¦ GPIO 5 (SCL) ¦ 400 kHz fast mode |
| ADC (ECG) | GPIO 0 ¦ 12-bit SAR ADC, 250 Hz via hardware timer |
| Flash | 4 MB |

## 2.2  BioAmp EXG Pill  –  ECG Acquisition

The BioAmp EXG Pill (Upside Down Labs) is a compact, single-supply bio-potential amplifier capable of capturing Electrocardiography (ECG), Electromyography (EMG), and Electroencephalography (EEG) signals. For AyuScan, it is configured for single-lead ECG (Lead I/II/III depending on electrode placement).

Gain: ~1000× differential amplification with common-mode rejection.

Bandwidth: 0.5 Hz – 40 Hz (suitable for QRS complex, P-wave, T-wave capture).

Output: Analog voltage signal sampled by ESP32-C6 ADC at 250 Hz (4 ms interrupt period).

Filtering: Hardware RC anti-aliasing + software 3rd-order Butterworth bandpass (0.5–40 Hz) in Python.

Electrode: 3 standard wet ECG electrodes (Right Arm, Left Arm, Right Leg reference).

## 2.3  MAX30102  –  PPG & SpO₂ Sensing

The Maxim Integrated MAX30102 is a pulse oximetry and heart rate sensor module featuring two integrated LEDs (Red: 660 nm, IR: 880 nm), a photodetector, optics, and low-noise analog signal processing. It communicates via I²C and stores data in an internal 32-sample FIFO.

| Parameter | Configuration in AyuScan |
| --- | --- |
| Red LED Pulse Amplitude | 0x3F (~12 mA) |
| IR LED Pulse Amplitude | 0x3F (~12 mA) |
| Green LED | Disabled (0x00) |
| Sampling Rate | 100 Hz (FIFO polled every loop iteration) |
| I²C Address | 0x57 |
| I²C Bus Speed | 400 kHz (fast mode) |

## 2.4  MLX90614  –  Infrared Body Temperature

The Melexis MLX90614 is a factory-calibrated infrared thermometer for non-contact temperature measurement. It uses a built-in 17-bit ADC and digital signal processor. In AyuScan, it is polled at 1 Hz. Since MLX90614 uses SMBus with Packet Error Checking (PEC), the I²C clock is temporarily stepped down to 100 kHz for every read operation, then restored to 400 kHz.

Object temperature range: −70°C to +380°C  |  Medical range: 32°C–42.5°C for body surface.

Accuracy: ±0.5°C for temperatures 0–50°C (body range).

I²C Address: 0x5A  |  SMBus PEC must be handled at 100 kHz.

Dual output: Object (body surface) and Ambient temperature reported separately in JSON payload.

## 2.5  Bone Conduction & Contact-Based PPG Principle

AyuScan leverages contact-based optical sensing, which shares the core physical principle with bone conduction technology: both rely on detecting biological signals transmitted through body tissue rather than through air. Understanding this relationship is critical to interpreting the PPG waveform quality.

### What is Bone Conduction?

Bone conduction is the transmission of sound or vibrational energy directly through the bones and cartilage of the skull or body, bypassing the external ear canal and tympanic membrane. In hearing aids and headphones, a transducer is pressed against the skull to vibrate the cochlea directly. In biomedical sensing, 'bone conduction' refers more broadly to the transmission of physiological vibrations — such as blood pulse pressure waves — directly through solid tissue.

### PPG as a Contact-Based Bio-Signal

Photoplethysmography (PPG) operates on the principle that blood absorbs infrared and red light differently depending on its oxygenation state and the volume of blood in the illuminated tissue. The MAX30102 shines its IR (880 nm) and Red (660 nm) LEDs into the fingertip or earlobe. Oxygenated hemoglobin (HbO₂) absorbs red light less and IR more, while deoxygenated hemoglobin (Hb) has the opposite property.

The PPG waveform has two primary components:

DC Component: Steady-state light absorption by venous blood, tissue, bone, and skin pigmentation.

AC Component: Pulsatile variation due to arterial blood volume changes with each heartbeat (systole/diastole).

The light passes through or reflects off tissue, traveling partly through bone and cartilage. This contact-based transmission is analogous to bone conduction — the signal is conducted through solid biological material, not air. The bone and tissue act as a waveguide for both the optical signal and the mechanical pulse pressure wave.

### SpO₂ Calculation (Ratio of Ratios)

AyuScan implements the industry-standard Maxim ratio-of-ratios algorithm for SpO₂ estimation:

R = (AC_red / DC_red) / (AC_ir / DC_ir)
SpO₂ = -45.060 × R² + 30.354 × R + 94.845

Where AC represents the pulsatile component and DC represents the mean (non-pulsatile) component for each LED channel. This empirical calibration curve maps the R-value to an SpO₂ percentage. Readings are smoothed using an Exponential Moving Average (α = 0.2) with a spike-rejection guard that discards values deviating more than 6% from the median history.

### Pulse Transit Time (PTT) – The Bone Conduction Analogy

Pulse Transit Time (PTT) is the time delay between an ECG R-peak (electrical cardiac event) and the arrival of the corresponding pulse pressure wave at a peripheral measurement site (fingertip PPG). This pressure wave travels through the arterial wall and surrounding tissue — including bone — via mechanical conduction, directly analogous to bone conduction of sound.

PTT is inversely correlated with blood pressure: stiffer arteries (higher BP) propagate the wave faster, resulting in shorter PTT. AyuScan measures PTT by detecting ECG R-peaks and the subsequent PPG foot (diastolic point) using scipy's find_peaks, with a physiological validity window of 100–500 ms.

# 3  Firmware  –  ayuScan_direct.ino

The firmware runs on the ESP32-C6 and is written in Arduino C++. It initialises all sensors, sets up a BLE GATT server with two notification characteristics, and uses a hardware timer ISR to achieve deterministic 250 Hz ECG sampling.

## 3.1  BLE Communication Protocol

| Element | Value |
| --- | --- |
| Service UUID | 4fafc201-1fb5-459e-8fcc-c5c9c331914b |
| ECG Characteristic UUID | beb5483e-36e1-4688-b7f5-ea07361b26a8 |
| PPG/Raw Characteristic UUID | e3223119-9445-4e96-a4a1-85358ce291d0 |
| MTU Size | 512 bytes |
| Property | NOTIFY (server pushes data to client without polling) |
| Device Name | AyuScan_Node |

JSON Packet Formats sent from device to host:

ECG Packet (100 ms batch, 25 samples at 250 Hz):

{"device":"patient_01","type":"ecg","fs":250,"data":[2048,2052,...,2045]}

PPG Raw Packet (100 ms batch, 10 samples at 100 Hz):

{"device":"patient_01","type":"ppg_raw","fs":100,"red":[45231,...],"ir":[67342,...]}

Temperature Packet (1 Hz):

{"device":"patient_01","type":"temp","val":36.72,"ambient":28.50}

## 3.2  Sensor Initialisation & I²C Bus Management

Both MAX30102 and MLX90614 share the same I²C bus. Since MLX90614 uses SMBus PEC which is unreliable above 100 kHz, the firmware implements a Hydrocheck-pattern: the clock is lowered to 100 kHz for every temperature read, then restored to 400 kHz immediately after.

Wire.begin(I2C_SDA=4, I2C_SCL=5)
Wire.setClock(400000)   // Fast mode for MAX30102
// ... MLX read ...
Wire.setClock(100000)   // SMBus mode for MLX90614
float temp = mlx.readObjectTempC()
Wire.setClock(400000)   // Restore fast mode

## 3.3  Timer-Driven ECG Acquisition at 250 Hz

A hardware timer (timerBegin at 1 MHz base clock) generates an interrupt every 4,000 ticks (4 ms = 250 Hz). The ISR sets a volatile flag ecgSampleReady using a port critical section. The main loop reads the ADC and batches 25 samples, then sends a BLE notification packet.

ecgTimer = timerBegin(1000000)      // 1 MHz base
timerAlarm(ecgTimer, 4000, true, 0) // 4ms period → 250 Hz
// ISR:
void IRAM_ATTR onTimer() { ecgSampleReady = true; }
// Main loop:
ecgBuffer[ecgIndex++] = analogRead(EXG_PIN);  // GPIO0

## 3.4  PPG FIFO Streaming at 100 Hz

The MAX30102 FIFO is polled continuously in the main loop using maxSensor.check(). Each available sample (IR + Red ADC value) is read and appended to a rolling buffer. When 10 samples accumulate (100 ms), a JSON PPG batch is sent over the PPG characteristic.

# 4  Backend  –  Python Gateway (ble_server.py)

The Python backend runs on the host PC and acts as a bridge between the BLE device and the web dashboard. It uses bleak for async BLE scanning/connection and websockets for the WS server. All AI models are loaded at startup and inference runs inline in the BLE notification callback.

## 4.1  BLE-to-WebSocket Bridge

ble_server.py runs two concurrent asyncio tasks: a WebSocket server (port 8080) that fans out all messages to connected browser clients, and a BLE worker that scans for 'AyuScan_Node', subscribes to both GATT characteristics, and processes every incoming notification.

Concurrency Model asyncio.gather(ws_server, ble_worker()) — both run on the same event loop. BLE notifications are async callbacks → they schedule asyncio tasks for broadcast. AI inference is synchronous (torch.no_grad) — kept fast enough not to block the loop.

## 4.2  PPG Signal Processing Algorithm (ppg_algorithm.py)

The PPGProcessor class implements a production-quality real-time signal processing pipeline. It maintains rolling buffers for BPM computation and SpO₂ estimation, and computes HRV metrics.

| Processing Stage | Description |
| --- | --- |
| IR Buffer (500 samples @ 100 Hz = 5 s) | Rolling ring buffer for systolic peak detection. |
| Bandpass Filter (0.75–2.5 Hz) | 2nd-order Butterworth removes baseline wander and high-frequency noise, passing 45–150 BPM. |
| Peak Detection | scipy.signal.find_peaks with distance=50 samples (0.5 s min), prominence=0.3×std, width=8 samples to reject dicrotic notch. |
| RR Interval Extraction | Inter-peak intervals in seconds; physiological gate: 0.5–1.2 s (50–120 BPM). |
| BPM Smoothing | Moving average over 8 BPM readings; last-good-BPM hold for up to 10 cycles. |
| HRV – SDNN | Standard deviation of all RR intervals in the last 30 cycles. |
| HRV – RMSSD | Root mean square of successive RR differences. |
| SpO₂ (25 Hz downsampled) | Ratio-of-ratios with EMA (α=0.2) and spike-rejection guard (±6% from median). |

## 4.3  Data Logging & Persistence

All acquired data is appended in real-time to two master CSV files:

datasets/ecg_data_master.csv  –  columns: timestamp, device, ecg_value (27 MB+ with continuous recording)

datasets/vitals_data_master.csv  –  columns: timestamp, device, spo2, hrv_sdnn, hrv_rmssd (500 KB+)

# 5  AI Model 1  –  ECG Cardiac Condition Classifier  (1D-CNN)

The ECG model performs real-time classification of ECG rhythm into four clinically relevant categories. It is implemented as a 1-Dimensional Convolutional Neural Network (1D-CNN) in PyTorch, specifically designed for temporal physiological signals.

## 5.1  Model Architecture  (ECG1DCNN)

Input: [Batch × 1 Channel × 1000 time steps]  (10 seconds @ 100 Hz, resampled from 250 Hz)

| Layer | Config | Output Shape |
| --- | --- | --- |
| Conv1D + BN + ReLU + MaxPool | in=1, out=16, kernel=7, pad=3 → Pool(2) | [B, 16, 500] |
| Conv1D + BN + ReLU + MaxPool | in=16, out=32, kernel=5, pad=2 → Pool(2) | [B, 32, 250] |
| Conv1D + BN + ReLU + MaxPool | in=32, out=64, kernel=5, pad=2 → Pool(2) | [B, 64, 125] |
| Conv1D + BN + ReLU + MaxPool | in=64, out=128, kernel=3, pad=1 → Pool(2) | [B, 128, 62] |
| AdaptiveAvgPool1D | Output size = 10 | [B, 128, 10] |
| Flatten | — | [B, 1280] |
| FC + ReLU + Dropout(0.5) | 1280 → 256 | [B, 256] |
| FC (output) | 256 → 4 classes | [B, 4] |

Total parameters: ~1.48 MB (ecg_model.pth)

Output classes:

Class 0: Normal – Regular sinus rhythm

Class 1: Bradycardia – Heart rate < 60 BPM

Class 2: Tachycardia – Heart rate > 100 BPM

Class 3: Arrhythmia – Irregular rhythm (AF, flutter, etc.)

## 5.2  Dataset  –  PTB-XL

PTB-XL is the world's largest openly available clinical 12-lead ECG dataset, released by Physikalisch-Technische Bundesanstalt (PTB) in collaboration with Berlin's Charité hospital.

| Attribute | Detail |
| --- | --- |
| Full Name | PTB-XL: A Large Publicly Available ECG Dataset |
| Records | 21,799 clinical 12-lead ECG recordings |
| Patients | 18,869 unique patients |
| Duration | 10 seconds per record |
| Sampling Rates | 500 Hz (high-res) and 100 Hz (low-res ← AyuScan uses lr) |
| Annotation | SCP-ECG codes (human-readable cardiovascular diagnosis codes) |
| Format | WFDB (WaveForm DataBase format, .hea + .dat) |
| License | Creative Commons Attribution 4.0 (CC BY 4.0) |
| Source | https://physionet.org/content/ptb-xl/ |
| Lead Used | Lead III (index 2 of 12) — approximates single-lead wrist ECG |

SCP Code Mapping used in AyuScan:

NORM: → Class 0 (Normal)

SBRAD: → Class 1 (Bradycardia)

STACH, SVTAC, PSVT: → Class 2 (Tachycardia)

AFIB, AFLT, SARRH, SVARR, BIGU, TRIGU: → Class 3 (Arrhythmia)

## 5.3  Training Procedure

| Hyperparameter | Value |
| --- | --- |
| Optimizer | Adam ¦ lr = 0.001 |
| Loss Function | CrossEntropyLoss |
| Epochs | 5 (prototype) ¦ Recommended: 20+ |
| Batch Size | 64 |
| Train/Val Split | 80% / 20% (random seed = 42) |
| Device | CUDA if available, else CPU |
| Preprocessing | WFDB read → Lead III extract → Butterworth 0.5–40 Hz → Z-score normalisation |
| Model Selection | Best checkpoint by minimum validation loss |

## 5.4  Inference Pipeline

At runtime, inference is triggered every time 2,500 ECG samples (10 s @ 250 Hz) accumulate:

1. Take last 2,500 raw ADC samples from ecg_buffers[device_id]

2. Resample from 2,500 → 1,000 points (scipy.signal.resample) to match training resolution

3. Apply 3rd-order Butterworth bandpass (0.5–40 Hz)

4. Z-score normalise (subtract mean, divide by std)

5. Shape into tensor [1, 1, 1000] and pass through ECG1DCNN

6. argmax of logits → class index → DISEASE_CLASSES lookup

7. Broadcast JSON {"type":"ai_prediction","condition":"Arrhythmia"} to all WS clients

# 6  AI Model 2  –  Blood Pressure Estimator  (XGBoost)

Cuffless blood pressure estimation is one of the most challenging problems in wearable health. AyuScan uses a dual XGBoost gradient boosting regressor — one for Systolic Blood Pressure (SBP) and one for Diastolic Blood Pressure (DBP) — trained on features derived from simultaneous ECG and PPG recordings with arterial line (ART) ground truth.

## 6.1  Feature Engineering

Five physiological features are extracted from every 5-second ECG+PPG window:

| Feature | Formula / Description | Clinical Significance |
| --- | --- | --- |
| hr | 60 / mean(RR intervals) | Heart rate in BPM from R-peak timing |
| ptt_mean | mean(R-peak → PPG foot delay) in seconds | Mean Pulse Transit Time — BP inversely correlated |
| ptt_std | std(PTT values) | PTT variability — reflects arterial compliance |
| ppg_amp_mean | mean(PPG peak − PPG foot) amplitude | Stroke volume proxy — related to DBP |
| rr_std | std(RR intervals) in seconds | HRV marker — related to sympathetic tone |

## 6.2  Dataset  –  VitalDB

VitalDB is a high-fidelity ICU waveform database collected from 6,388 surgical patients at Seoul National University Hospital. AyuScan uses the downloadable CSV exports of perioperative vital sign waveforms with simultaneous invasive arterial blood pressure (ART), ECG Lead II, and PPG (PLETH) channels.

| Attribute | Detail |
| --- | --- |
| Full Name | VitalDB: A High-Fidelity Perioperative Vital Sign Database |
| Cases | 6,388 surgical patients |
| Channels Used | ECG_II, PLETH (PPG), ART (invasive arterial BP) |
| Sampling Rate | 100 Hz (all channels synchronised) |
| Ground Truth BP | Invasive arterial line (gold standard — no calibration error) |
| Window Size | 5 seconds (500 samples) with 1-second step |
| Max Samples | 5,000 windows (fast extraction mode) |
| License | Open access for research |
| Source | https://vitaldb.net/ |

## 6.3  Model Details & Metrics

| Parameter | SBP Model | DBP Model |
| --- | --- | --- |
| Algorithm | XGBoost XGBRegressor | XGBoost XGBRegressor |
| n_estimators | 100 | 100 |
| learning_rate | 0.1 | 0.1 |
| max_depth | 5 | 5 |
| tree_method | hist (GPU-accelerated histogram) | hist |
| Train/Test Split | 80% / 20% (seed = 42) | 80% / 20% (seed = 42) |
| Target Range | 70–220 mmHg (SBP clipped) | 40–130 mmHg (DBP clipped) |

AHA Blood Pressure Classification (applied to predictions):

| Category | SBP (mmHg) | DBP (mmHg) |
| --- | --- | --- |
| Normal | < 120 | < 80 |
| Elevated | 120–129 | < 80 |
| Hypertension Stage 1 | 130–139 | 80–89 |
| Hypertension Stage 2 | ≥ 140 | ≥ 90 |
| Hypertensive Crisis | > 180 | > 120 |

# 7  AI Model 3  –  SpO₂ Trend Classifier  (GRU)

Unlike the other models that classify instantaneous measurements, the SpO₂ model performs temporal trend analysis — predicting whether a patient's oxygen saturation will remain stable or decline in the near future. This makes it a proactive early-warning system, not a reactive alert.

## 7.1  Model Architecture  (SpO2_GRU)

Input: [Batch × 60 time steps × 1 feature] — 60 consecutive SpO₂ readings (approx. 60 seconds)

| Layer | Details |
| --- | --- |
| GRU Layer | input_size=1, hidden_size=16, num_layers=1, batch_first=True |
| Output extraction | Last hidden state output[:, -1, :] → shape [B, 16] |
| FC Layer | 16 → 4 classes |
| Total parameters | ~6.9 KB (spo2_model.pth) — extremely lightweight for edge deployment |

Alternative CNN model (SpO2_CNN) is also defined and can be selected at training time.

## 7.2  Dataset  –  BIDMC Multiparameter Dataset

The Beth Israel Deaconess Medical Center (BIDMC) PPG and Respiration Dataset provides simultaneous PPG and SpO₂ recordings from critically ill ICU patients. AyuScan uses the numeric CSV exports (bidmc_csv/XX_Numerics.csv).

| Attribute | Detail |
| --- | --- |
| Full Name | BIDMC PPG and Respiration Dataset |
| Subjects | 53 ICU patients |
| Duration | 8 minutes each |
| SpO₂ Column | ' SpO2' (note leading space in CSV) sampled at ~1 Hz |
| Window Size | 60 samples (60 seconds look-back) |
| Future Window | 30 samples (30 seconds look-ahead for label generation) |
| License | Open Database License (ODC-ODbL) |
| Source | https://physionet.org/content/bidmc/ |

## 7.3  Label Schema & Inference

Labels are generated by a sliding-window predictive labelling algorithm:

| Class | Label | Condition |
| --- | --- | --- |
| 0 | Stable | SpO₂ ≥ 90% and future drop < 1.5% |
| 1 | Mild Decline | Future minimum drops 1.5–2.9% below current |
| 2 | Rapid Decline | Future minimum drops ≥ 3% below current |
| 3 | Critical | Current or future SpO₂ < 90% |

Inference is gated: SpO₂ AI runs only after 300 valid readings (≈5 minutes) are collected, with a 5-minute cooldown between alerts to prevent alert fatigue. Normalisation: (SpO₂ − 90) / 10.

Training Configuration:

Epochs: 20  |  Batch size: 32  |  Optimizer: Adam lr=0.001  |  Loss: CrossEntropyLoss

Class balancing: Majority class (Stable) undersampled to max 3× minority count

Evaluation: Classification report + confusion matrix + per-sample CPU inference time

# 8  Dashboard  –  React Frontend

The AyuScan Monitor dashboard is a Vite-powered React SPA that connects to the WebSocket server and displays real-time ECG waveforms, vital signs, and AI predictions for multiple patients simultaneously.

Key UI features:

Multi-patient grid layout: 1, 2, or 4 cards per page; auto-carousel every 10 s for ≥ 5 devices.

PatientCard component: Live ECG chart (canvas/SVG), BPM, SpO₂, temperature, HRV, AI condition badge.

Device Assignment Modal: Assign a BLE device ID and patient name.

Sidebar: Quick-access patient list with device status.

PatientDetailsModal: Full-screen detailed view of a selected patient.

WebSocket hook (useLiveDevices): Manages device state, historical data buffers, and WS reconnection.

Tech stack: React 18, Vite, Lucide icons, Vanilla CSS with CSS custom properties (design tokens).

# 9  System Data Flow

1.  Hardware Sensors: ECG @ 250 Hz via ADC  |  PPG IR+Red @ 100 Hz via MAX30102 FIFO  |  Temp @ 1 Hz via MLX90614

↓

2.  ESP32-C6 Firmware: Buffers samples, encodes JSON, sends BLE NOTIFY to ble_server.py

↓

3.  BLE Gateway: bleak receives notifications, decodes JSON, logs to CSV, routes by type

↓

4.  PPG Processing: PPGProcessor: bandpass → peak detection → BPM, SpO₂, HRV

↓

5.  ECG AI Inference: 2500 samples → resample → filter → normalise → 1D-CNN → condition class

↓

6.  BP Inference: 500 samples each ECG+PPG → extract 5 features → XGBoost → SBP/DBP mmHg

↓

7.  SpO₂ AI Inference: 60 valid SpO₂ readings → normalize → GRU → trend class (gated, cooldown)

↓

8.  WebSocket Broadcast: JSON messages pushed to all connected browser clients

↓

9.  React Dashboard: useLiveDevices hook updates device state → PatientCard re-renders with live data

# 10  Dependencies & Environment

Python Backend Requirements:

| Package | Version | Purpose |
| --- | --- | --- |
| bleak | 0.21+ | Async BLE client for Windows/Linux/macOS |
| websockets | 12.0+ | Async WebSocket server |
| torch (PyTorch) | 2.0+ | 1D-CNN and GRU model inference |
| numpy | 1.24+ | Numerical signal processing |
| scipy | 1.10+ | Butterworth filters, peak detection, resampling |
| xgboost | 1.7+ | Blood pressure regression models |
| joblib | 1.2+ | Model serialisation (XGBoost .pkl files) |
| pandas | 2.0+ | Data loading and feature extraction |
| wfdb | 4.0+ | Reading PTB-XL WFDB waveform files |
| scikit-learn | 1.2+ | Preprocessing, metrics, train/test split |
| tqdm | 4.65+ | Progress bars for dataset extraction |

Frontend Requirements (Node.js / npm):

React 18  |  Vite 5  |  Lucide React icons

Hardware Requirements:

Seeed Studio XIAO ESP32-C6  |  BioAmp EXG Pill  |  MAX30102 module  |  MLX90614 module

3× ECG electrodes (disposable Ag/AgCl)  |  Jumper wires  |  5V USB power bank

# 11  Deployment Guide

## Step 1: Flash Firmware

1. Open Arduino IDE → Board Manager → Install 'esp32' by Espressif
2. Select Board: 'XIAO_ESP32C6'
3. Install libraries: SparkFun MAX3010x, Adafruit MLX90614, ESP32 BLE Arduino
4. Open ayuScan_direct/ayuScan_direct.ino
5. Upload to device via USB

## Step 2: Setup Python Backend

cd d:/AyuScan
pip install -r requirements.txt
python ble_server.py

## Step 3: Train AI Models (optional — pre-trained .pth files included)

# ECG Model
cd backend/AI-Model
python train.py

# SpO2 Model
cd backend/Spo2-Mocdel
python data_pipeline.py   # generates spo2_dataset.npz
python train.py

# BP Model
cd backend/Bp-model
python download_vitaldb.py   # downloads dataset CSVs
python extract_dataset_fast.py   # generates extracted_features.csv
python train.py

## Step 4: Start Dashboard

cd dashboard
npm install
npm run dev
# Open http://localhost:5173

# 12  Clinical Safety Notes

⚠ NOT a Certified Medical Device AyuScan has NOT been tested or approved by FDA, CE, or any regulatory body. Do NOT use AyuScan for clinical decision-making, diagnosis, or treatment. SpO₂ readings may be less accurate on patients with dark skin pigmentation (known limitation of PPG oximetry). Blood pressure estimation via PTT requires individual calibration against a cuff measurement for accuracy. ECG classification accuracy is bounded by the training dataset distribution; rare arrhythmias may be misclassified. Always seek professional medical advice for any health concerns.

# 13  Glossary

1D-CNN:  1-Dimensional Convolutional Neural Network — applies convolutional filters along the time axis of a signal.

AC/DC (PPG):  Alternating/Direct Current components of the PPG signal corresponding to pulsatile and non-pulsatile light absorption.

BLE:  Bluetooth Low Energy — a power-efficient wireless protocol for short-range data transmission.

ECG:  Electrocardiogram — electrical recording of cardiac activity via surface electrodes.

GRU:  Gated Recurrent Unit — a recurrent neural network variant that models temporal dependencies with reset and update gates.

HRV:  Heart Rate Variability — variation in time between consecutive heartbeats; a marker of autonomic nervous system health.

PPG:  Photoplethysmography — optical measurement of blood volume changes in tissue using light.

PTT:  Pulse Transit Time — time for a cardiac pulse pressure wave to travel from the heart to a peripheral site.

RMSSD:  Root Mean Square of Successive Differences in RR intervals — short-term HRV measure, parasympathetic tone proxy.

SDNN:  Standard Deviation of Normal-to-Normal RR intervals — overall HRV measure.

SpO₂:  Peripheral Oxygen Saturation — percentage of haemoglobin carrying oxygen, estimated non-invasively via PPG.

SBP / DBP:  Systolic / Diastolic Blood Pressure — peak and trough arterial pressure during the cardiac cycle (mmHg).

WFDB:  WaveForm DataBase — an open standard for storing and reading physiological waveform data.

XGBoost:  Extreme Gradient Boosting — a highly efficient ensemble learning algorithm based on gradient-boosted decision trees.

AyuScan Technical Documentation  |  Generated 26 August 2026, 10:08  |  For Research Use Only

AyuScan — Quick Summary

🔧 Hardware

- ESP32-C6 (RISC-V, 160 MHz) — BLE GATT server, ADC, I²C controller

- BioAmp EXG Pill — ECG acquisition, ~1000× amplification, 0.5–40 Hz bandwidth

- MAX30102 — IR (880nm) + Red (660nm) PPG sensor, 100 Hz, I²C 0x57

- MLX90614 — Non-contact IR thermometer, 1 Hz, I²C 0x5A (SMBus 100 kHz)

📡 Firmware (ayuScan_direct.ino)

- Hardware timer ISR at 250 Hz for ECG (GPIO 0 ADC)

- PPG polled from MAX30102 FIFO at 100 Hz

- JSON packets sent via BLE NOTIFY in 100ms batches

- Two GATT characteristics: ECG + PPG/Temp

🦴 Bone Conduction / PPG Principle

- PPG light passes through tissue & bone (contact-based conduction)

- AC component = pulsatile blood volume → BPM, SpO₂

- DC component = tissue/bone/venous baseline

- PTT (Pulse Transit Time) = pressure wave conducted through arterial wall/tissue → inversely correlates with BP

🤖 AI Model 1 — ECG Classifier (1D-CNN, PyTorch)

- 4 conv layers (16→32→64→128 filters) + AdaptiveAvgPool + FC

- Input: 1000 time steps (10s @ 100 Hz after resampling)

- Dataset: PTB-XL — 21,799 clinical ECG records (100 Hz low-res, Lead III)

- Classes: Normal | Bradycardia | Tachycardia | Arrhythmia

- Model size: ~1.48 MB | Loss: CrossEntropyLoss | Optimizer: Adam

🤖 AI Model 2 — Blood Pressure Estimator (XGBoost)

- Features: HR, PTT mean, PTT std, PPG amplitude mean, RR std

- Dataset: VitalDB — 6,388 ICU patients, invasive ART ground truth, 100 Hz

- Dual regressors: SBP model + DBP model (100 trees, depth=5)

- Outputs AHA category: Normal → Elevated → HTN Stage 1/2 → Crisis

🤖 AI Model 3 — SpO₂ Trend Classifier (GRU, PyTorch)

- GRU: hidden=16, input=1, → FC → 4 classes

- Input: 60 consecutive SpO₂ readings (~60 seconds)

- Dataset: BIDMC — 53 ICU patients, ~1 Hz SpO₂

- Classes: Stable | Mild Decline | Rapid Decline | Critical

- Gated: runs only after 300 valid readings, 5-min cooldown

- Model size: only 6.9 KB

🖥️ Dashboard (React + Vite)

- WebSocket (port 8080) → real-time updates

- Multi-patient grid (1/2/4 cards), auto-carousel for 5+ devices

- Shows: live ECG waveform, BPM, SpO₂, HRV (SDNN/RMSSD), Temp, BP, AI badge

📊 Datasets Summary

| Dataset | Used For | Size |
| --- | --- | --- |
| PTB-XL | ECG classifier | 21,799 records |
| VitalDB | BP regression | 6,388 patients |
| BIDMC | SpO₂ trend | 53 ICU patients |